中科院地质与地球物理研究所

MC-ICP-MS

实验室分析方法描述

2025.12.29 北京

为了方便大家写论文时正确描述实验室的分析方法及仪器参数，中科院地质与地球物理研究所MC-ICP-MS实验室特提供以下分析方法描述范例，以供参考 （2025.12）。

注：本文件为MC-ICP-MS实验室分析方法的笼统性描述，请各位用户根据实验过程中采用的分析仪器参考相应的分析方法描述，并根据具体实验条件（文中红色字体部分）进行修改.

目录

1 四级杆(Q)-ICP-MS	3

1.1 In situ trace element analysis by LA-Q-ICP-MS	3

1.2 Zircon U-Pb dating and trace element analyses by LA-Q-ICP-MS	4

1.3 In situ bastnäsite trace element and U-Th-Pb age analysis	5

1.4 In situ titanite trace element and U-Pb age analysis	7

1.5 In situ monazite trace element and U-Pb age analysis	8

1.6 In situ xenotime trace element and U-Pb age analysis	10

1.7 In situ columbite-tantalite trace element and U-Pb age analysis	11

1.8 In situ allanite trace element and U-Pb age analysis	13

1.9 In situ perovskite trace element and U-Pb age analysis	14

2 扇形磁场双聚焦(SF)-ICP-MS	16

2.1 In situ analysis of ultra-low trace elements by LA-SF-ICP-MS	16

2.2 Concurrent determinations of Fo value and trace element of olivine by LA-SF-ICP-MS	17

2.3 High spatial resolution zircon U-Pb dating by LA-SF-ICP-MS	17

2.4 In situ wolframite U-Pb dating by LA-SF-ICP-MS	18

2.5 In situ garnet U-Pb dating by LA-SF-ICP-MS	20

2.6 In situ vesuvianite U-Pb dating by LA-SF-ICP-MS	21

2.7 In situ cassiterite U-Pb dating by LA-SF-ICP-MS	23

2.8 In situ calcite U-Pb dating by LA-SF-ICP-MS	24

3 多接收(MC)-ICP-MS	26

3.1 In situ Hf isotope ratio analysis by LA-MC-ICP-MS	26

3.2 In situ Sr isotope analysis by LA-MC-ICP-MS	28

3.3 In situ Nd isotope analysis by LA-MC-ICP-MS	31

3.4 Sr-Nd-Hf-Pb isotope analysis by solution MC-ICP-MS	36

1 四级杆(Q)-ICP-MS

1.1 In situ trace element analysis by LA-Q-ICP-MS

方法描述

Trace element contents of clinopyroxene /amphibole… were determined by laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an iCAP TQ ICP-Q-MS/MS coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The approach is similar to those outlined in Xie et al. (2008) with isotopes measured using a peak-hopping mode with a laser beam diameter of ca. XX m and XX Hz repetition rate. The laser energy density is ~4.0 J/cm2. Helium was employed as the ablation gas to improve the transporting efficiency of ablated aerosols. ARM-1 reference glass (Wu et al., 2019, 2021) was used as calibration material, and NIST SRM 610 (Jochum et al., 2011) and BCR-2G (Jochum et al., 2005) were analyzed for data quality control. Silicon (29Si) was used as an internal standard. The resulting data were reduced based on the GLITTER program (Griffin et al., 2008). For most trace elements (>0.10 g/g), the accuracy is better than ±10 % with analytical precision (1 RSD) of ±10%.

参考文献

Griffin, W., Powell, W., Pearson, N.J., O Reilly, S., 2008. GLITTER: data reduction software for laser ablation ICP-MS. In: Sylvester, P. (Eds.), Laser Ablation-ICP-MS in the Earth Sciences: Current Practices and Outstanding Issues, pp. 308-311.

Jochum, K.P., Willbold, M., Raczek, I., Stoll, B., Herwig, K., 2005. Chemical characterisation of the USGS reference glasses GSA-1G, GSC-1G, GSD-1G, GSE-1G, BCR-2G, BHVO-2G and BIR-1G using EPMA, ID-TIMS, ID-ICP-MS and LA-ICP-MS. Geostandards and Geoanalytical Research 29, 285-302.

Jochum, K.P., Weis, U., Stoll, B., Kuzmin, D., Yang, Q., Raczek, I., Jacob, D.E., Stracke, A., Birbaum, K., Frick, D.A., Günther, D., Enzweiler, J., 2011. Determination of reference values for NIST SRM 610-617 glasses following ISO guidelines. Geostandards and Geoanalytical Research 35, 397-429.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.(这篇文章报道了ARM玻璃的主微量数据，请用户根据实际情况选择引用文献)

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719-745. (这篇文章报道了ARM玻璃的主量、同位素以及Fe2+/ΣFe比值数据，请用户根据实际情况选择引用文献)

Xie, L.W., Zhang, Y.B., Zhang, H.H., Sun, J.F., Wu, F.Y., 2008. In situ simultaneous determination of trace elements, U-Pb and Lu-Hf isotopes in zircon and baddeleyite. Chinese Science Bulletin 53, 1565-1573.

1.2 Zircon U-Pb dating and trace element analyses by LA-Q-ICP-MS

方法描述

The U-Pb isotope compositions of zircon were measured in situ using an iCAP TQ ICP-Q-MS/MS coupled with a GeoLasHD 193 nm ArF excimer laser ablation system or an Analyte G2 193 nm ArF excimer laser ablation system at Institute of Geology and Geophysics, Chinese Academy of Sciences. Detailed analytical procedure followed those outlined in Xie et al. (2008). For most data reported here the density of energy at the ablation spots was ~4.0 J cm-2 at 80 mJ output of energy and laser beam diameters were ca. XXm at a repetition rate of XX Hz. The dwell time for each isotope was set at 6 ms for 29Si, 49Ti, 93Nb, 181Ta, 91Zr and REE, 15, 15, 30, 15 ms for 204Pb, 206Pb, 207Pb and 208Pb and 10 ms for 232Th and 238U. Each spot analysis comprised 20 s of gas background followed by 50 s of sample ablation. ARM-1 (Wu et al., 2019, 2021) was applied to calibrate the trace element concentration for all spot analyses. 91500 zircon (TIMS 207Pb/206Pb age = 1065.4 ± 0.3 Ma, 1s, n = 11; Wiedenbeck et al. 1995) or GJ-1 zircon (TIMS 206Pb/238U age = 599.8 ± 4.5 Ma; Jackson et al., 2004) was treated as an external reference zircon to calibrate mass bias and instrument drift, which measured twice every ten analyses. Another reference zircon SA01 (TIMS 206Pb/238U age = 535.08 ± 0.32 Ma; Huang et al., 2020) was analyzed as an unknown to monitor the quality of age data. The 207Pb/206Pb, 206Pb/238U, 207Pb/235U and 208Pb/232Th ratios were calculated using the GLITTER program, whereas the 235U signal was calculated from 238U on the basis of the ratio 238U/235U = 137.818 (Hiess et al. 2012). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard (SA01) during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties of 207Pb/206Pb and 206Pb/238U were estimated to be ~0.5% and ~1.5%, respectively, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties. The age calculations and plotting of Concordia diagrams were made using ISOPLOT (v3.23) (Ludwig 2003).

参考文献

Griffin, W., Powell, W., Pearson, N.J., O Reilly, S., 2008. GLITTER: data reduction software for laser ablation ICP-MS. In: Sylvester, P. (Eds.), Laser Ablation-ICP-MS in the Earth Sciences: Current Practices and Outstanding Issues, pp. 308-311.

Hiess, J., Condon, D.J., McLean, N., Noble, S.R., 2012. 238U/235U Systematics in Terrestrial Uranium-Bearing Minerals. Science 335, 1610-1614.

Huang, C., Wang, H., Yang, J., Ramezani, J., Yang, C., Zhang, S., Yang, Y., Xia, X., Feng, L., Lin, J., Wang, T., Ma, Q., He, H., Xie, L., Wu, S., 2020. SA01-A proposed zircon reference material for microbeam U-Pb age and Hf-O isotopic determination. Geostandards and Geoanalytical Research 44, 103-123.

Jackson, S.E., Pearson, N.J., Griffin, W.L., Belousova, E.A., 2004. The application of laser ablation-inductively coupled plasma-mass spectrometry to in situ U-Pb zircon geochronology. Chemical Geology 211, 47-69.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Wiedenbeck, M., Alle, P., Corfu, F., Griffin, W.L., Meier, M., Oberli, F., Vonquadt, A., Roddick, J.C., Speigel, W., 1995. Three natural zircon standards for U-Th-Pb, Lu-Hf, trace element and REE analyses. Geostandards Newsletter 19, 1-23.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719-745.

Xie, L.W., Zhang, Y.B., Zhang, H.H., Sun, J.F., Wu, F.Y., 2008. In situ simultaneous determination of trace elements, U-Pb and Lu-Hf isotopes in zircon and baddeleyite. Chinese Science Bulletin 53, 1565-1573.

1.3 In situ bastnäsite trace element and U-Th-Pb age analysis

方法描述

An iCAP TQ ICP-Q-MS/MS coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system was used to determine trace element compositions and U-Pb ages. A matrix-matched in house reference material consisting of a crystal of K-9 bastnäsite with TIMS U-Pb and Th-Pb age of 116.80 ± 0.13 Ma and 116.59 ± 0.11 Ma (Sal'Nikova et al., 2010; Li et al., 2020) was used to correct for U-Th-Pb fractionation and instrumental mass discrimination. Two K-9 and one LZ1384/MAD809 (Yang et al., 2019) or TPZ18Q (Qu et al., 2019) analyses were measured after every five unknown bastnäsite spots. More details about the analytical methods were provided in Yang et al. (2014, 2019). The 207Pb correction method was applied for common Pb correction assuming a Pb composition equivalent to Stacey and Kramers (1975), and the intercept ages or 206Pb/238U weighted ages were calculated using Isoplot 3.23 (Ludwig et al., 2003). Moreover, the intercepts of the regression line through the raw data on a Tera-Wasserburg plot provide an estimate of the 207Pb/206Pb for the common Pb component (upper intercept) and the inferred crystallization age (lower intercept). Similarly, considering the lower U content of most bastnäsites, we also conducted 208Pb/232Th age calculation after common 207Pb correction (Yang et al., 2019). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard (LZ1384 or MAD809 or TPZ18Q) during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties of the U-Pb ages were estimated to be ~2 %, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties.

Quantitative results for trace elements were obtained via the external calibration of relative elemental sensitivities using ARM-1 standard (Wu et al., 2019, 2021) and normalization of each analysis to 140Ce as the internal standard using the Glitter software. The Ce2O3 content in bastnäsite samples was determined by electron microprobe analysis.

参考文献

Li, Q.L., Huyskens, M.H., Yang, Y.H., Ling, X.X., Yin, Q.Z., Nikiforov, A., Li, X.H., 2020. Bastnaesite K-9: A Homogenous Natural Reference Material for In-situ U-Pb and Th-Pb Dating. Atomic Spectroscopy 41, 218-222.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Qu, K., Sima, X., Zhou, H., Xiao, Z., Tu, J., Yin, Q., Liu, X., Li, J., 2019. In situ LA-MC-ICP-MS and ID-TIMS U-Pb ages of bastnäsite-(Ce) and zircon from the Taipingzhen hydrothermal REE deposit: New constraints on the later Paleozoic granite-related U-REE mineralization in the North Qinling Orogen, Central China. Journal of Asian Earth Sciences 173, 352-363.

Sal'Nikova, E.B., Yakovleva, S.Z., Nikiforov, A.V., Kotov, A.B., Yarmolyuk, V.V., Anisimova, I.V., Sugorakova, A.M., Plotkina, Y.V., 2010. Bastnaesite: a promising U-Pb geochronological tool. Doklady Earth Sciences 430, 134-136.

Stacey, J.S., Kramers, J.D., 1975. Approximation of terrestrial lead isotope evolution by a two-stage model. Earth and Planetary Science Letters 26, 207-221.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

Yang, Y., Wu, F., Li, Q., Rojas-Agramonte, Y., Yang, J., Li, Y., Ma, Q., Xie, L., Huang, C., Fan, H., Zhao, Z., Xu, C., 2019. In Situ U-Th-Pb Dating and Sr-Nd Isotope Analysis of Bastnasite by LA-(MC)-ICP-MS. Geostandards and Geoanalytical Research 43, 543-565.

Yang, Y., Wu, F., Li, Y., Yang, J., Xie, L., Liu, Y., Zhang, Y., Huang, C., 2014. In situ U-Pb dating of bastnaesite by LA-ICP-MS. Journal of Analytical Atomic Spectrometry 29, 1017-1023.

1.4 In situ titanite trace element and U-Pb age analysis

方法描述

An iCAP TQ ICP-Q-MS/MS coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system was used to determine trace element compositions and U-Pb ages. A matrix matched titanite reference material (BLR-1) (Aleinikoff et al., 2007) was used as external standard to correct 207Pb/206Pb, 206Pb/238U, 207Pb/235U (235U = 238U/137.818) and 208Pb/232Th ratios. Two BLR-1 and one Ontraio or OLT-1 analyses were measured after every five to eight unknown samples. All measured 207Pb/206Pb, 207Pb/235U and 206Pb/238U isotopic ratios were regressed and corrected following the method of Sun et al. (2012) and Ma et al. (2019). Trace element concentrations were calibrated using 43Ca as the internal standard (CaO contents were measured previously by EPMA) and using ARM-1 (Wu et al., 2019, 2021) as the external standard material. Isotopic and elemental fractionation plus instrumental mass bias were calibrated using Glitter 4.0 software (Griffin et al. 2008). For multiple groups of standards, we select the option for the interpolation of linear fit to ratios to perform drift corrections. Signal sections of each analysis were selected independently to get the very similar interval for standards and unknowns. The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard (Ontraio) (Spencer et al., 2013) or OLT-1 (Kennedy et al., 2010) during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. The U-Pb ages and weighted mean ages were calculated using the ISOPLOT 3.0 software package (Ludwig et al., 2003). Systematic uncertainties were estimated to be ~2.0%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties.

参考文献

Aleinikoff, J.N., Wintsch, R.P., Tollo, R.P., Unruh, D.M., Fanning, C.M., Schmitz, M.D., 2007. Ages and origins of rocks of the Killingworth dome, south-central Connecticut: Implications for the tectonic evolution of southern New England. American Journal of Science 307, 63-118.

Griffin, W., Powell, W., Pearson, N.J., O Reilly, S., 2008. GLITTER: data reduction software for laser ablation ICP-MS. In: Sylvester, P. (Eds.), Laser Ablation-ICP-MS in the Earth Sciences: Current Practices and Outstanding Issues, pp. 308-311.

Kennedy, A.K., Kamo, S.L., Nasdala, L., Timms, N.E., 2010. Grenville skarn titanite: potential reference material for SIMS U-Th-Pb analysis. The Canadian Mineralogist 48, 1423-1443.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Ma, Q., Evans, N.J., Ling, X., Yang, J., Wu, F., Zhao, Z., Yang, Y., 2019. Natural Titanite Reference Materials for In Situ U-Pb and Sm-Nd Isotopic Measurements by LA-(MC)-ICP-MS. Geostandards and Geoanalytical Research 43, 355-384.

Spencer, K.J., Hacker, B.R., Kylander-Clark, A.R.C., Andersen, T.B., Cottle, J.M., Stearns, M.A., Poletti, J.E., Seward, G.G.E., 2013. Campaign-style titanite U–Pb dating by laser-ablation ICP: Implications for crustal flow, phase transformations and titanite closure. Chemical Geology 341, 84-101.

Sun, J., Yang, J., Wu, F., Xie, L., Yang, Y., Liu, Z., Li, X., 2012. In situ U-Pb dating of titanite by LA-ICPMS. Chinese Science Bulletin 57, 2506-2516.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

1.5 In situ monazite trace element and U-Pb age analysis

方法描述

An iCAP TQ ICP-Q-MS/MS coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system was used to determine trace element compositions and U-Pb ages. A matrix matched monazite reference material (44069 (Aleinikoff et al., 2006) or RW-1 (Ling et al., 2017)) was used as external standard to correct 207Pb/206Pb, 206Pb/238U, 207Pb/235U, and 208Pb/232Th ratios. Monazite reference materials (Diamantina (Gonçalves et al., 2018) or Vermilion (Mazoz et al., 2024)) were used as quality control materials. Two 44069 (or RW-1) and one Diamantina (or Vermilion) analyses were measured after every eight unknown samples. All measured 207Pb/206Pb, 206Pb/238U, 207Pb/235U and 208Pb/232Th isotopic ratios were regressed and corrected following the method of Liu et al. (2012). Trace element concentrations were calibrated using 140Ce as the internal standard (Ce2O3 contents were measured previously by EPMA) and using ARM-1 (Wu et al., 2019, 2021) as the external standard material. Isotopic and elemental fractionation plus instrumental mass bias were calibrated using Glitter 4.0 software (Griffin et al. 2008). For multiple groups of standards, we select the option for the interpolation of linear fit to ratios to perform drift corrections. Signal sections of each analysis were selected independently to get the very similar interval for standards and unknowns. The U-Pb and Th-Pb ages and weighted mean ages were calculated using the ISOPLOT 3.0 software package (Ludwig et al., 2003). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~1.5%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties.

参考文献

Aleinikoff, J.N., Schenck, W.S., Plank, M.O., Srogi, L.A., Fanning, C.M., Kamo, S.L., Bosbyshell, H., 2006. Deciphering igneous and metamorphic events in high-grade rocks of the Wilmington Complex, Delaware: Morphology, cathodoluminescence and backscattered electron zoning, and SHRIMP U-Pb geochronology of zircon and monazite. Geological Society of America Bulletin 118, 39-64.

Gonçalves, G.O., Lana, C., Scholz, R., Buick, I.S., Gerdes, A., Kamo, S.L., Corfu, F., Rubatto, D., Wiedenbeck, M., Nalini Jr., H.A., Oliveira, L.C.A., 2018. The Diamantina Monazite: A New Low-Th Reference Material for Microanalysis. Geostandards and Geoanalytical Research 42, 25-47.

Griffin, W., Powell, W., Pearson, N.J., O Reilly, S., 2008. GLITTER: data reduction software for laser ablation ICP-MS. In: Sylvester, P. (Eds.), Laser Ablation-ICP-MS in the Earth Sciences: Current Practices and Outstanding Issues, pp. 308-311.

Ling, X.X., Huyskens, M.H., Li, Q.L., Yin, Q.Z., Werner, R., Liu, Y., Tang, G.Q., Yang, Y.N., Li, X.H., 2017. Monazite RW-1: a homogenous natural reference material for SIMS U-Pb and Th-Pb isotopic analysis. Mineralogy and Petrology 111, 163-172.

Liu, Z., Wu, F., Yang, Y., Yang, J., Wilde, S. 2012. Neodymium isotopic compositions of the standard monazites used in U-Th-Pb geochronology. Chemical Geology 334: 221-239.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Mazoz, A., Gonçalves, G.O., Lana, C., Buick, I.S., McSwiggen, P.L., Corfu, F., Kamo, S.L., Fu, B., Wang, H., Moreira, H., Scholz, R., Martins, L., Peixoto, E., Dantas, E.L., Santos, R.V., 2024. Vermilion monazite: A new Archean reference material for U-Pb and Sm-Nd microanalysis. Chemical Geology 658, 122121.

Sun, J., Yang, J., Wu, F., Xie, L., Yang, Y., Liu, Z., Li, X., 2012. In situ U-Pb dating of titanite by LA-ICPMS. Chinese Science Bulletin 57, 2506-2516.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

1.6 In situ xenotime trace element and U-Pb age analysis

方法描述

An iCAP TQ ICP-Q-MS/MS coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system was used to determine trace element compositions and U-Pb ages. A matrix matched xenotime reference material (MG-1) was used as external standard to correct 207Pb/206Pb, 207Pb/235U, and 206Pb/238U ratios, and xenotime reference material BS-1 was used as quality control material (Fletcher et al., 2004). Two MG-1 and one BS-1 analyses were measured after every eight unknown samples. All measured 207Pb/206Pb, 207Pb/235U, and 206Pb/238U isotopic ratios were regressed and corrected following the method of Liu et al. (2011). Trace element concentrations were calibrated using 89Y as the internal standard (Y2O3 contents were measured previously by EPMA) and using ARM-1 (Wu et al., 2019, 2021) as the external standard material. Isotopic and elemental fractionation plus instrumental mass bias were calibrated using Glitter 4.0 software (Griffin et al. 2008). For multiple groups of standards, we select the option for the interpolation of linear fit to ratios to perform drift corrections. Signal sections of each analysis were selected independently to get the very similar interval for standards and unknowns. The U-Pb ages and weighted mean ages were calculated using the ISOPLOT 3.0 software package (Ludwig et al., 2003). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~1.5%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties.

参考文献

Fletcher, I.R., McNaughton, N.J., Aleinikoff, J.A., Rasmussen, B., Kamo, S.L., 2004. Improved calibration procedures and new standards for U-Pb and Th-Pb dating of Phanerozoic xenotime by ion microprobe. Chemical Geology 209, 295-314.

Griffin, W., Powell, W., Pearson, N.J., O Reilly, S., 2008. GLITTER: data reduction software for laser ablation ICP-MS. In: Sylvester, P. (Eds.), Laser Ablation-ICP-MS in the Earth Sciences: Current Practices and Outstanding Issues, pp. 308-311.

Liu, Z.C., Wu, F. Y., Guo, C. L., Zhao, Z. F., Yang, J. H., Sun, J. F. 2011. In situ U-Pb dating of xenotime by laser ablation (LA)-ICP-MS. Chinese Science Bulletin, 56: 2948-2956.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

1.7 In situ columbite-tantalite trace element and U-Pb age analysis

方法描述

An iCAP TQ ICP-Q-MS/MS coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system was used to determine trace element compositions and U-Pb ages. A matrix matched columbite-tantalite reference material (Coltan139, 507.9±1.3 Ma) (Melcher et al., 2015; Yang et al., 2023) was used as external standard to correct 207Pb/206Pb, 207Pb/235U, and 206Pb/238U ratios, while columbite-tantalite reference material ZKW (Yang et al., 2023) or SN3 (Xiang et al., 2023) was used as quality control material. Two Coltan139 and one ZKW/SN3 analyses were measured after every eight unknown samples. All measured 207Pb/206Pb, 207Pb/235U and 206Pb/238U isotopic ratios during sample analyses were regressed and corrected following the method of Che et al. (2015) and Yang et al., (2023). Trace element concentrations were calibrated using 93Nb or 55Mn as the internal standard (Nb2O5 or MnO contents were measured previously by EPMA) and using ARM-1 (Wu et al., 2019, 2021) as the external standard material. Isotopic and elemental fractionation plus instrumental mass bias were calibrated using Glitter 4.0 software (Griffin et al. 2008). For multiple groups of standards, we select the option for the interpolation of linear fit to ratios to perform drift corrections. Signal sections of each analysis were selected independently to get the very similar interval for standards and unknowns. The U-Pb ages and weighted mean ages were calculated using the ISOPLOT 3.0 software package (Ludwig et al., 2003). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~2.0%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties was propagated to the final reported uncertainties.

参考文献

Che, X-D., Wu, F-Y., Wang, R-C., Gerdes, A., Ji, W-Q., Zhao, Z-H., Yang, J-H., Zhu, Z-Y., 2015. In situ U-Pb isotopic dating of columbite-tantalite by LA-ICP-MS. Ore Geology Reviews. 65: 979-989.

Griffin, W., Powell, W., Pearson, N.J., O Reilly, S., 2008. GLITTER: data reduction software for laser ablation ICP-MS. In: Sylvester, P. (Eds.), Laser Ablation-ICP-MS in the Earth Sciences: Current Practices and Outstanding Issues, pp. 308-311.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Melcher, F., Graupner, T., Gäbler, H., Sitnikova, M., Henjes-Kunst, F., Oberthür, T., Gerdes, A., Dewaele, S., 2015. Tantalum-(niobium-tin) mineralisation in African pegmatites and rare metal granites: Constraints from Ta-Nb oxide mineralogy, geochemistry and U-Pb geochronology. Ore Geology Reviews 64, 667-719.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

Xiang, L., Wang, R., Romer, R.L., Che, X., Hu, H., Tang, Z., 2023. Columbite SN3: A new potential reference material for U-Pb dating by LA-ICP-MS. Geostandards and Geoanalytical Research 47, 609-628.

1.8 In situ allanite trace element and U-Pb age analysis

方法描述

In situ allanite U-Pb dating was performed using laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an iCAP TQ ICP-Q-MS/MS (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The details of measurement procedures and calibration technique can be found in Yang et al., (2022). Laser diameter of 30-44 m and 5 Hz repetition rate and energy density of ~3.0 J/cm2 were used. Allanite standard LE40010 (or Tara) is used as the primary calibration material, while standard allanite CAPb (or LE2808 or A007) (Yang et al., 2022) is analyzed as a quality control material. The 207Pb correction method was applied for common Pb correction assuming a Pb composition equivalent to Stacey and Kramers (1975). The intercept ages or 206Pb/238U weighted ages were calculated using Isoplot 3.23 (Ludwig et al., 2003). Moreover, the intercepts of the regression line through the raw data on a Tera-Wasserburg plot provide an estimate of the 207Pb/206Pb for the common Pb component (upper intercept) and the inferred crystallization age (lower intercept). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~2.0%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties. Trace element mass fractions were calibrated using Si as the internal standard element (SiO2 contents were measured previously by EPMA) and using NIST SRM 610 (Jochum et al. 2011) as the primary reference material and ARM-3 (Wu et al., 2019, 2021) for data monitoring.

参考文献

Jochum, K.P., Weis, U., Stoll, B., Kuzmin, D., Yang, Q., Raczek, I., Jacob, D.E., Stracke, A., Birbaum, K., Frick, D.A., Günther, D., Enzweiler, J., 2011. Determination of reference values for NIST SRM 610-617 glasses following ISO guidelines. Geostandards and Geoanalytical Research 35, 397-429.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Stacey, J.S., Kramers, J.D., 1975. Approximation of terrestrial lead isotope evolution by a two-stage model. Earth and Planetary Science Letters 26, 207-221.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

Yang, M., Yang, Y. H., Kamo, S. L., Romer, R. L., Roberts, N. M., Wang, H., Xie L., W., Huang C., Yang J., H., & Wu, F. Y. (2022). Natural Allanite Reference Materials for In Situ U-Th-Pb and Sm-Nd Isotopic Measurements by LA-(MC)-ICP-MS. Geostandards and Geoanalytical Research, 46(2), 169-203.

1.9 In situ perovskite trace element and U-Pb age analysis

方法描述

In situ perovskite U-Pb dating was performed using laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an iCAP TQ ICP-Q-MS/MS (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The details of measurement procedures and calibration technique can be found in Yang et al., (2009). Laser diameter of 44 m and 4 Hz repetition rate and energy density of ~3.0 J/cm2 were used. The Tazheran perovskite (Kinny et al., 1997) is used as the primary calibration material, while in-house standard AFK-5 (Wu et al., 2013) is analyzed as a quality control material. The 207Pb correction method was applied for common Pb correction assuming a Pb composition equivalent to Stacey and Kramers (1975). The intercept ages or 206Pb/238U weighted ages were calculated using Isoplot 3.23 (Ludwig et al., 2003). Moreover, the intercepts of the regression line through the raw data on a Tera-Wasserburg plot provide an estimate of the 207Pb/206Pb for the common Pb component (upper intercept) and the inferred crystallization age (lower intercept). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~2.0%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties. Trace element mass fractions were calibrated using Ca as the internal standard element (CaO contents were measured previously by EPMA) and using NIST SRM 610 (Jochum et al. 2011) as the primary reference material and ARM-3 (Wu et al., 2019, 2021) for data monitoring.

参考文献

Jochum, K.P., Weis, U., Stoll, B., Kuzmin, D., Yang, Q., Raczek, I., Jacob, D.E., Stracke, A., Birbaum, K., Frick, D.A., Günther, D., Enzweiler, J., 2011. Determination of reference values for NIST SRM 610-617 glasses following ISO guidelines. Geostandards and Geoanalytical Research 35, 397-429.

Kinny, P.D., Griffin, B.J., Heaman, L.M., Brakhfogel, F.F., Spetsius, Z.V., 1997. SHRIMP U-Pb ages of perovskite from Yakutian kimberlites. Russian Geology and Geophysics 38, 97-105.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Stacey, J.S., Kramers, J.D., 1975. Approximation of terrestrial lead isotope evolution by a two-stage model. Earth and Planetary Science Letters 26, 207-221.

Wu, F., Arzamastsev, A.A., Mitchell, R.H., Li, Q., Sun, J., Yang, Y., Wang, R., 2013. Emplacement age and Sr-Nd isotopic compositions of the Afrikanda alkaline ultramafic complex, Kola Peninsula, Russia. Chemical Geology 353, 210-229.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

Yang, Y.H., Wu, F.Y., Wilde, S.A., Liu, X.M., Zhang, Y.B., Xie, L.W., Yang, J.H., 2009. In situ perovskite Sr-Nd isotopic constraints on the petrogenesis of the Ordovician Mengyin kimberlites in the North China Craton. Chemical Geology 264, 24-42.

2 扇形磁场双聚焦(SF)-ICP-MS

2.1 In situ analysis of ultra-low trace elements by LA-SF-ICP-MS

方法描述

Trace element abundances of olivine/orthopyroxene/carbonate minerals… were determined by laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an Element XR HR-ICP-MS instrument (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The approach is similar to that outlined in Wu et al. (2019) with isotopes measured using a peak-hopping mode with a laser diameter of ca. XX m and XX Hz repetition rate. The laser energy density is ~4.0 J/cm2. The Element XR is equipped with the so-called “jet-interface”, comprising of a jet sample cone, an X-version skimmer cone and a high capacity vacuum pump (OnTool Booster 150, Asslar, Germany). This leads to a signal enhancement in laser sampling mode by a factor of 3-5, resulting in an improved detection capability. Helium was employed as the ablation gas to improve the transporting efficiency of ablated aerosols. ARM-3 reference glass was used for external calibration (Wu et al., 2019, 2021) and BIR-1G glass (Jochum et al., 2005) was used for quality control monitoring. Silicon (29Si) or Calcium (43Ca) was used as an internal standard. The resulting data were reduced based on the GLITTER program (Griffin et al., 2008). For most trace elements (>0.005 g/g), the accuracy is better than ±10 % with analytical precision (1 RSD) of ±10%.

参考文献

Griffin, W., Powell, W., Pearson, N.J., O Reilly, S., 2008. GLITTER: data reduction software for laser ablation ICP-MS. In: Sylvester, P. (Eds.), Laser Ablation-ICP-MS in the Earth Sciences: Current Practices and Outstanding Issues, pp. 308-311.

Jochum, K.P., Willbold, M., Raczek, I., Stoll, B., Herwig, K., 2005. Chemical characterisation of the USGS reference glasses GSA-1G, GSC-1G, GSD-1G, GSE-1G, BCR-2G, BHVO-2G and BIR-1G using EPMA, ID-TIMS, ID-ICP-MS and LA-ICP-MS. Geostandards and Geoanalytical Research 29, 285-302.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

2.2 Concurrent determinations of Fo value and trace element of olivine by LA-SF-ICP-MS

方法描述

Simultaneous quantification of forsterite (Fo) content and minor-trace elements in olivine was performed by laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an Element XR HR-ICP-MS instrument (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. Laser diameter of ca. 44 m and 5 Hz repetition rate and energy density of ~4.0 J/cm2 were used. The measured isotopes are 7Li, 23Na, 25Mg, 27Al, 29Si, 31P, 43Ca, 45Sc, 49Ti, 51V, 53Cr, 55Mn, 57Fe, 59Co, 60Ni, 63Cu, 67Zn, 69Ga and 89Y. The details of measurement procedures and calibration technique can be found in Wu et al., (2020). The resulting data were reduced based on the Iolite program (Paton et al. 2011) with a home-made DRS approach (Wu et al. 2018). The precision and accuracy of Fo values are better than 0.3. The detection limits for most trace elements are better than 0.05 g/g.

参考文献

Paton, C., Hellstrom, J., Paul, B., Woodhead, J., Hergt, J., 2011. Iolite: Freeware for the visualisation and processing of mass spectrometric data. Journal of Analytical Atomic Spectrometry 26, 2508-2518.

Wu, S., Wu, Y., Yang, Y., Wang, H., Huang, C., Xie, L., Yang, J., 2020. Simultaneous Quantification of Forsterite Content and Minor–Trace Elements in Olivine by LA–ICP–MS and Geological Applications in Emeishan Large Igneous Province. Minerals 10, 634; doi:10.3390/min10070634.

Wu, S., Huang, C., Xie, L., Yang, Y., Yang, J., 2018. Iolite Based Bulk Normalization as 100% (m/m) Quantification Strategy for Reduction of Laser Ablation-Inductively Coupled Plasma-Mass Spectrometry Transient Signal. Chinese Journal of Analytical Chemistry 46, 1628-1636.

2.3 High spatial resolution zircon U-Pb dating by LA-SF-ICP-MS

方法描述

High spatial resolution (5-16 m) zircon U-Pb dating was performed using laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an Element XR HR-ICP-MS instrument (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The details of measurement procedures and calibration technique can be found in Wu et al., (2020). Laser diameter of 5-16 m and 5 Hz repetition rate and energy density of ~3.0 J/cm2 were used. The Element XR is equipped with the so-called “jet-interface”, comprising of a jet sample cone, an X-version skimmer cone and a high capacity vacuum pump (OnTool Booster 150, Asslar, Germany). This leads to a signal enhancement in laser sampling mode by a factor of 3-5, resulting in an improved detection capability. Helium was employed as the ablation gas to improve the transporting efficiency of ablated aerosols. Standard zircon 91500 (Wiedenbeck et al. 1995) is used as the primary calibration material. Newly developed standard zircon SA01 (Huang et al., 2020) is analyzed as a quality control material. The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard (SA01) during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~2.0 %, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. This uncertainty was propagated to the final reported uncertainties.

参考文献

Huang, C., Wang, H., Yang, J., Ramezani, J., Yang, C., Zhang, S., Yang, Y., Xia, X., Feng, L., Lin, J., Wang, T., Ma, Q., He, H., Xie, L., Wu, S., 2020. SA01-A proposed zircon reference material for microbeam U-Pb age and Hf-O isotopic determination. Geostandards and Geoanalytical Research 44, 103-123.

Wiedenbeck, M., Alle, P., Corfu, F., Griffin, W.L., Meier, M., Oberli, F., Vonquadt, A., Roddick, J.C., Speigel, W., 1995. Three natural zircon standards for U-Th-Pb, Lu-Hf, trace element and REE analyses. Geostandards Newsletter 19, 1-23.

Wu, S., Yang, M., Yang, Y., Xie, L., Huang, C., Wang, H., Yang, J., 2020. Improved in situ zircon U-Pb dating at high spatial resolution (5-16 μm) by laser ablation-single collector-sector field-ICP-MS using Jet sample and X skimmer cones. International Journal of Mass Spectrometry 456, 116394.

2.4 In situ wolframite U-Pb dating by LA-SF-ICP-MS

方法描述

High spatial resolution (30-44 m) wolframite U-Pb dating was performed using laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an Element XR HR-ICP-MS instrument (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The details of measurement procedures and calibration technique can be found in Yang et al., (2020). Laser diameter of 30-44 m and 5 Hz repetition rate and energy density of ~3.0 J/cm2 were used. Newly developed in house wolframite YGX-2113 is used as the primary calibration material, while newly developed standard wolframite Sewa (or SHM or YGX-2107) (Yang et al., 2020) is analyzed as a quality control material. The 207Pb correction method was applied for common Pb correction assuming a Pb composition equivalent to Stacey and Kramers (1975), intercept ages or 206Pb/238U weighted ages were calculated using Isoplot 3.23 (Ludwig et al., 2003). Moreover, the intercepts of the regression line through the raw data on a Tera-Wasserburg plot provide an estimate of the 207Pb/206Pb for the common Pb component (upper intercept) and the inferred crystallization age (lower intercept). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~2.5%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties. The concentrations of U-Th-Pb were reduced using Iolite software with “Trace_Element” DRS and the semi-quantitative standardisation method (Paton et al. 2011). NIST SRM 612 (Jochum et al. 2011) is used as external calibrator and ARM-3 (Wu et al., 2019, 2021) is used for the quality control. The yielded results match with the published data with a discrepancy of 20 %.

参考文献

Jochum, K.P., Weis, U., Stoll, B., Kuzmin, D., Yang, Q., Raczek, I., Jacob, D.E., Stracke, A., Birbaum, K., Frick, D.A., Günther, D., Enzweiler, J., 2011. Determination of reference values for NIST SRM 610-617 glasses following ISO guidelines. Geostandards and Geoanalytical Research 35, 397-429.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Paton, C., Hellstrom, J., Paul, B., Woodhead, J., Hergt, J., 2011. Iolite: Freeware for the visualisation and processing of mass spectrometric data. Journal of Analytical Atomic Spectrometry 26, 2508-2518.

Stacey, J.S., Kramers, J.D., 1975. Approximation of terrestrial lead isotope evolution by a two-stage model. Earth and Planetary Science Letters 26, 207-221.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

Yang, M., Yang, Y., Wu, S., Romer, R.L., Che, X., Zhao, Z., Li, W., Yang, J., Wu, F., Xie, L., Huang, C., Zhang, D., Zhang, Y., 2020. Accurate and precise in situ U-Pb isotope dating of wolframite series minerals via LA-SF-ICP-MS. Journal of Analytical Atomic Spectrometry 35, 2191-2203.

2.5 In situ garnet U-Pb dating by LA-SF-ICP-MS

方法描述

High spatial resolution (30-44 m) garnet (schorlomite and grossular-andradite) U-Pb dating was performed using laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an Element XR HR-ICP-MS instrument (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The details of measurement procedures and calibration technique can be found in Yang et al. (2025). Laser diameter of 30-44 m and 5 Hz repetition rate and energy density of ~3.0 J/cm2 were used. Newly developed in house garnet PL34 (or Willsboro) is used as the primary calibration material, while newly developed standard garnet MKWB (or Mali) (Yang et al., 2025; Seman et al., 2017) is analyzed as a quality control material. The 207Pb correction method was applied for common Pb correction assuming a Pb composition equivalent to Stacey and Kramers (1975), intercept ages or 206Pb/238U weighted ages were calculated using Isoplot 3.23 (Ludwig et al., 2003). Moreover, the intercepts of the regression line through the raw data on a Tera-Wasserburg plot provide an estimate of the 207Pb/206Pb for the common Pb component (upper intercept) and the inferred crystallization age (lower intercept). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~2.5%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties. The concentrations of U-Th-Pb were reduced using Iolite software with “Trace_Element” DRS and the semi-quantitative standardisation method (Paton et al. 2011). NIST SRM 612 (Jochum et al. 2011) is used as external calibrator and ARM-3 (Wu et al., 2019, 2021) is used for the quality control. The yielded results match with the published data with a discrepancy of 20 %.

参考文献

Jochum, K.P., Weis, U., Stoll, B., Kuzmin, D., Yang, Q., Raczek, I., Jacob, D.E., Stracke, A., Birbaum, K., Frick, D.A., Günther, D., Enzweiler, J., 2011. Determination of reference values for NIST SRM 610-617 glasses following ISO guidelines. Geostandards and Geoanalytical Research 35, 397-429.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Paton, C., Hellstrom, J., Paul, B., Woodhead, J., Hergt, J., 2011. Iolite: Freeware for the visualisation and processing of mass spectrometric data. Journal of Analytical Atomic Spectrometry 26, 2508-2518.

Seman, S., Stockli, D.F., McLean, N.M., 2017. U-Pb geochronology of grossular-andradite garnet. Chemical Geology 460, 106-116.

Stacey, J.S., Kramers, J.D., 1975. Approximation of terrestrial lead isotope evolution by a two-stage model. Earth and Planetary Science Letters 26, 207-221.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

Yang, Y.H., Wu, S.T., Wang, H., Kamo, S.L., Ma, Q., Liang, T., Xu, L., Xie, L.W., Huang, C., Wan, B., Yang, J.H., Wu, F.Y., 2025. Three new natural secondary reference materials for in situ andradite U-Pb geochronology. Journal of Analytical Atomic Spectrometry 40, 326-337.

2.6 In situ vesuvianite U-Pb dating by LA-SF-ICP-MS

方法描述

High spatial resolution (60 m) vesuvianite U-Pb dating was performed using laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an Element XR HR-ICP-MS instrument (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The details of measurement procedures and calibration technique can be found in Wei et al., (2022). Laser diameter of 60 m and 5 Hz repetition rate and energy density of ~5.0 J/cm2 were used. Garnet standard PL34 (or vesuvianite standard Wilui) was used as the primary reference material, respectively, while newly developed in house vesuvianite Wilui (or M6635) was used as the quality control material (Wei et al. 2022). The 207Pb correction method was applied for common Pb correction assuming a Pb composition equivalent to Stacey and Kramers (1975), intercept ages or 206Pb/238U weighted ages were calculated using Isoplot 3.23 (Ludwig et al., 2003). Moreover, the intercepts of the regression line through the raw data on a Tera-Wasserburg plot provide an estimate of the 207Pb/206Pb for the common Pb component (upper intercept) and the inferred crystallization age (lower intercept). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~2.5 %, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties. The concentrations of U-Th-Pb were reduced using Iolite software with “Trace_Element” DRS and the semi-quantitative standardisation method (Paton et al. 2011). NIST SRM 612 (Jochum et al. 2011) is used as external calibrator and ARM-3 (Wu et al., 2019, 2021) is used for the quality control. The yielded results match with the published data with a discrepancy of 20 %.

参考文献

Jochum, K.P., Weis, U., Stoll, B., Kuzmin, D., Yang, Q., Raczek, I., Jacob, D.E., Stracke, A., Birbaum, K., Frick, D.A., Günther, D., Enzweiler, J., 2011. Determination of reference values for NIST SRM 610-617 glasses following ISO guidelines. Geostandards and Geoanalytical Research 35, 397-429.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Paton, C., Hellstrom, J., Paul, B., Woodhead, J., Hergt, J., 2011. Iolite: Freeware for the visualisation and processing of mass spectrometric data. Journal of Analytical Atomic Spectrometry 26, 2508-2518.

Stacey, J.S., Kramers, J.D., 1975. Approximation of terrestrial lead isotope evolution by a two-stage model. Earth and Planetary Science Letters 26, 207-221.

Wei, Q.D., Yang, M., Romer, R.L., Wang, H., Yang, Y.H., Zhao, Z.F., Wu, S.T., Xie, L.W., Huang, C., Xu, L., Yang, J.H., Wu, F.Y., 2022. In situ U-Pb geochronology of vesuvianite by LA-SF-ICP-MS. Journal of Analytical Atomic Spectrometry 37, 69-81.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

2.7 In situ cassiterite U-Pb dating by LA-SF-ICP-MS

方法描述

High spatial resolution (30-44 m) cassiterite U-Pb dating was performed using laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an Element XR HR-ICP-MS instrument (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The details of measurement procedures and calibration technique can be found in Yang et al., (2022). Laser diameter of 30-44 m and 5 Hz repetition rate and energy density of ~3.0 J/cm2 were used. Cassiterite reference material SPG-II (Rizvanova and Kuznetsov, 2020) is used as the primary calibration material, while newly developed standard cassiterite Rond-A (or PitAB) (Yang et al., 2022) is analyzed as a quality control material. The 207Pb correction method was applied for common Pb correction assuming a Pb composition equivalent to Stacey and Kramers (1975), intercept ages or 206Pb/238U weighted ages were calculated using Isoplot 3.23 (Ludwig et al., 2003). Moreover, the intercepts of the regression line through the raw data on a Tera-Wasserburg plot provide an estimate of the 207Pb/206Pb for the common Pb component (upper intercept) and the inferred crystallization age (lower intercept). The relative standard deviation (RSD) of the 206Pb/238U ratios obtained from the monitor standard during the same analytical sessions with the unknown samples was propagated to the uncertainties of each analytical spots. Systematic uncertainties were estimated to be ~2.5%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties.The concentrations of U-Th-Pb were reduced using Iolite software with “Trace_Element” DRS and the semi-quantitative standardisation method (Paton et al. 2011). NIST SRM 612 (Jochum et al. 2011) is used as external calibrator and ARM-3 (Wu et al., 2019, 2021) is used for the quality control. The yielded results match with the published data with a discrepancy of 20 %.

参考文献

Jochum, K.P., Weis, U., Stoll, B., Kuzmin, D., Yang, Q., Raczek, I., Jacob, D.E., Stracke, A., Birbaum, K., Frick, D.A., Günther, D., Enzweiler, J., 2011. Determination of reference values for NIST SRM 610-617 glasses following ISO guidelines. Geostandards and Geoanalytical Research 35, 397-429.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Paton, C., Hellstrom, J., Paul, B., Woodhead, J., Hergt, J., 2011. Iolite: Freeware for the visualisation and processing of mass spectrometric data. Journal of Analytical Atomic Spectrometry 26, 2508-2518.

Rizvanova, N.G., Kuznetsov, A.B., 2020. A New Approach to ID-TIMS U–Pb Dating of Cassiterite by the Example of the Pitkäranta Tin Deposit. Doklady Earth Sciences 491, 146-149.

Stacey, J.S., Kramers, J.D., 1975. Approximation of terrestrial lead isotope evolution by a two-stage model. Earth and Planetary Science Letters 26, 207-221.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

Yang, M., Romer, R. L., Yang, Y. H., Wu, S. T., Wang, H., Tu, J. R., Zhou, H., Y., Xie. L. W., Huang C., Xu L., Yang J. H., and Wu, F. Y. 2022. U-Pb isotopic dating of cassiterite: Development of reference materials and in situ applications by LA-SF-ICP-MS. Chemical Geology, 593, 120754.

2.8 In situ calcite U-Pb dating by LA-SF-ICP-MS

方法描述

High spatial resolution (85-100 m) calcite U-Pb dating was performed using laser ablation-inductively coupled plasma-mass spectrometry (LA-ICP-MS) employing an Element XR HR-ICP-MS instrument (Thermo Fisher Scientific, USA) coupled to a 193nm ArF excimer laser system (Geolas HD, Lambda Physik, Göttingen, Germany) or an Analyte G2 193 nm ArF excimer laser ablation system at the State Key Laboratory of Lithospheric and Environmental Coevolution, Institute of Geology and Geophysics, Chinese Academy of Sciences. The details of measurement procedures and calibration technique can be found in Wu et al., (2022). Laser diameter of 85-100 m and 10 Hz repetition rate and energy density of ~3.0 J/cm2 were used. WC-1 (Roberts et al., 2017) is used as the primary calibration material. Duff Brown Tank (Hill et al., 2016) is analyzed as a quality control material. Intercept ages calculated using Isoplot 3.23 (Ludwig et al., 2003). Moreover, the intercepts of the regression line through the raw data on a Tera-Wasserburg plot provide an estimate of the 207Pb/206Pb for the common Pb component (upper intercept) and the inferred crystallization age (lower intercept). Systematic uncertainties were estimated to be ~2.5%, which include the uncertainty of reference material, long-term variance of the validation material, uncertainty of the decay constant, and other potential sources. These uncertainties were propagated to the final reported uncertainties. The concentrations of U-Th-Pb were reduced using Iolite software with “Trace_Element” DRS and the semi-quantitative standardisation method (Paton et al. 2011). NIST SRM 614 (Jochum et al. 2011) is used as external calibrator and ARM-3 (Wu et al., 2019, 2021) is used for the quality control. The yielded results match with the published data with a discrepancy of 20 %.

参考文献

Hill, C. A., Polyak, V. J., Asmerom, Y., & P. Provencio, P. (2016). Constraints on a Late Cretaceous uplift, denudation, and incision of the Grand Canyon region, southwestern Colorado Plateau, USA, from U-Pb dating of lacustrine limestone. Tectonics, 35(4), 896-906.

Jochum, K.P., Weis, U., Stoll, B., Kuzmin, D., Yang, Q., Raczek, I., Jacob, D.E., Stracke, A., Birbaum, K., Frick, D.A., Günther, D., Enzweiler, J., 2011. Determination of reference values for NIST SRM 610-617 glasses following ISO guidelines. Geostandards and Geoanalytical Research 35, 397-429.

Ludwig, K.R., 2003. ISOPLOT 3.00: A Geochronological Toolkit for Microsoft Excel. Berkeley Geochronology Center, California, Berkeley.

Paton, C., Hellstrom, J., Paul, B., Woodhead, J., Hergt, J., 2011. Iolite: Freeware for the visualisation and processing of mass spectrometric data. Journal of Analytical Atomic Spectrometry 26, 2508-2518.

Roberts, N. M., Rasbury, E. T., Parrish, R. R., Smith, C. J., Horstwood, M. S., & Condon, D. J. (2017). A calcite reference material for LA-ICP-MS U-Pb geochronology. Geochemistry, Geophysics, Geosystems, 18(7), 2807-2814.

Wu, S., Worner, G., Jochum, K.P., Stoll, B., Simon, K., Kronz, A., 2019. The Preparation and Preliminary Characterisation of Three Synthetic Andesite Reference Glass Materials (ARM-1, ARM-2, ARM-3) for In Situ Microanalysis. Geostandards and Geoanalytical Research 43, 567-584.

Wu S., Yang Y., Jochum K.P., Romer R.L., Glodny J., Savov I.P., Agostini S., De Hoog J.C.M., Peters S.T.M., Kronz A., Zhang C., Bao Z., Wang X., Li Y., Tang G., Feng L., Yu H., Li Z., Le Z., Lin J., Zeng Y., Xu C., Wang Y., Cui Z., Deng L., Xiao J., Liu Y., Xue D., Di Z., Jia L., Wang H., Xu L., Huang C., Xie L., Pack A., Wörner G., He M., Li C., Yuan H., Huang F., Li Q., Yang J., Li X. and Wu F. 2021 Isotopic Compositions (Li-B-Si-O-Mg-Sr-Nd-Hf-Pb) and Fe2+/ΣFe Ratios of Three Synthetic Andesite Glass Reference Materials (ARM-1, ARM-2, ARM-3), Geostandards and Geoanalytical Research, 45, 719–745.

Wu, S., Yang, Y., Roberts, N. M., Yang, M., Wang, H., Lan, Z., ... & Wu, F. (2022). In situ calcite U-Pb geochronology by high-sensitivity single-collector LA-SF-ICP-MS. Science China Earth Sciences, 65(6), 1146-1160.

3 多接收(MC)-ICP-MS

3.1 In situ Hf isotope ratio analysis by LA-MC-ICP-MS

In situ Hf isotope ratio measurements were performed using a Neptune (Plus) multi-collector inductively coupled plasma mass spectrometer (MC-ICP-MS) (Thermo Fisher Scientific, USA). The analyses were conducted at the Institute of Geology and Geophysics, Chinese Academy of Sciences (IGGCAS), Beijing, China, using a GeoLasPro 193 nm ArF excimer laser (Coherent, Germany) or an Analyte G2 193 nm ArF excimer laser (PhotonMachines, USA) equipped with an ANU HelEx 2-volume cell. Detailed analytical procedures are provided in Wu et al. (2006), Xie et al. (2008), and Huang et al. (2022).

All analyses used a laser spot size of 44 μm (32 μm), a repetition rate of 6 Hz, and a fluence of 6.0 J cm-2. Each measurement consisted of 200 cycles with an integration time of 0.131s per cycle. To enhance sensitivity, nitrogen gas was additionally mixed into the carrier gas. The measured 176Hf/177Hf ratios were corrected for mass fractionation using an exponential law, with 179Hf/177Hf = 0.7325 (Patchett and Tatsumoto, 1980). Precise corrections for isobaric interferences from 176Lu and 176Yb on 176Hf were achieved using a 173Yb/172Yb ratio of 0.73925, a 176Yb/172Yb ratio of 0.5887, and 176Lu/175Lu ratio of 0.02655 (Vercoort et al., 2004) and assuming spot-specific βYb=βLu values.

Data quality was assessed by analyzing the zircon reference materials SA01 and SA02 /baddeleyite reference materials Phalaborwa and SK10-3 as unknowns. The resulting 176Hf/177Hf values of ___ ± __ (2s) and  ___ ± __ (2s) were in agreement, within analytical uncertainty, with the previously reported values of 0.282293 ± 0.000007 (2s) and 0.282287 ± 0.000016 (2s) (Huang et al., 2020, Huang et al., 2021), respectively/0.281229 ± 0.000011 (2s) and 0.282742 ± 0.000008 (2s) (Wu et al., 2006, Chen et al., 2025), obtained through solution-based methods.

如果数据进行了外部标样校正，则标样数据描述情况如下：

Zircon SA01 (Huang et al., 2021) was used as the primary reference material, while zircon SA02 (0.282287 ± 0.000016, 2s; Huang et al., 2021) and synthetic zircon MUNZirc 4 (0.282139 ± 0.000004, 2s; Fisher et al., 2011) served as monitoring reference materials for Hf isotopic ratio analysis. The measured 176Hf/177Hf ratios for both monitoring reference materials (___ ± __ , 2s and ___ ± __, 2s) were consistent with their recommended values within uncertainty. The uncertainty associated with zircon SA01 (0.282293 ± 0.000007, 2s; Huang et al., 2020) was propagated to the final uncertainty of the sample measurements.

Baddeleyite Phalaborwa (Wu et al., 2006) was used as the primary reference material, while SK10-3 (0.282742 ± 0.000008, 2s; Chen et al., 2025) served as monitoring reference materials for Hf isotopic ratio analysis. The measured 176Hf/177Hf ratio for the monitoring reference material (___ ± __ , 2s) was consistent with its recommended value within uncertainty. The uncertainty associated with Phalaborwa (0.281229 ± 0.000011, 2s; Wu et al., 2006) was propagated to the final uncertainty of the sample measurements.

参考文献

Chen, R, Wu, S., Wang, H., Senger, M., Paul, A. N, Paul, J. S., Yang, Y., Yang, J., Wu, F. 2025 Baddeleyite SK10-3: A Natural Reference Material for Microbeam U-Pb Geochronology and Hf Isotopic Measurements. Geostandards and Geoanalytical Research 49, 179-195.

Huang, C., Wang, H., Yang, J., Ramezani, J., Yang, C., Zhang, S., Yang, Y., Xia, X., Feng, L., Lin, J., Wang, T., Ma, Q., He, H., Xie, L., Wu, S., 2020. SA01-A proposed zircon reference material for microbeam U-Pb age and Hf-O isotopic determination. Geostandards and Geoanalytical Research 44, 103-123.

Huang, C., Yang, Y., Wang, H., Xie, L., Wu, S., Xu, L., ... & Wu, F. (2022). In-run measuring 177Hf16O/177Hf as a routine technique for in-situ Hf isotopic compositions analysis in zirconium-bearing minerals by laser ablation MC-ICP-MS. Spectrochimica Acta Part B: Atomic Spectroscopy, 194, 106486.

Huang, C., Wang, H., Yang, J., Yang, C., Li, Z., Yang, Y., Feng, L., Xie, L., Wu, S., 2021. Characterization of the potential reference material SA02 for micro-beam U–Pb geochronology and Hf–O isotopic composition analysis of zircon. Journal of Analytical Atomic Spectrometry 36, 368-374.

Patchett, P.J., Tatsumoto, M., 1980. A routine high-precision method for Lu-Hf isotope geochemistry and chronology. Contributions to Mineralogy and Petrology 75, 263-267.

Vervoort, J.D., Patchett, P.J., Soderlund, U., Baker, M., 2004. Isotopic composition of Yb and the determination of Lu concentrations and Lu/Hf ratios by isotope dilution using MC-ICPMS. Geochemistry Geophysics Geosystems 5. Q11002. https://doi.org/10.1029/2004GC000721.

Wu, F.Y., Yang, Y.H., Xie, L.W., Yang, J.H., Xu, P., 2006. Hf isotopic compositions of the standard zircons and baddeleyites used in U-Pb geochronology. Chemical Geology 234, 105-126.

Xie, L.W., Zhang, Y.B., Zhang, H.H., Sun, J.F., Wu, F.Y., 2008. In situ simultaneous determination of trace elements, U-Pb and Lu-Hf isotopes in zircon and baddeleyite. Chinese Science Bulletin 53, 1565-1573.

Fisher, C. M., Hanchar, J. M., Samson, S. D., Dhuime, B., Blichert-Toft, J., Vervoort, J. D., Lam, R., 2011. Synthetic zircon doped with hafnium and rare earth elements: A reference material for in situ hafnium isotope analysis. Chemical Geology 286, 32-47.

3.2 In situ Sr isotope analysis by LA-MC-ICP-MS

In situ Sr isotopic measurements were performed using a Neptune (Plus) multi-collector inductively coupled plasma mass spectrometer (MC-ICP-MS) (Thermo Fisher Scientific, USA). The analyses were conducted at the Institute of Geology and Geophysics, Chinese Academy of Sciences (IGGCAS), Beijing, China, using a GeoLasPro 193 nm ArF excimer laser (Coherent, Germany) or an Analyte G2 193 nm ArF excimer laser (PhotonMachines, USA) equipped with an ANU HelEx 2-volume cell. Detailed analytical procedures are provided in Yang et al. (2009a).

Laser spot sizes ranged from 40 to 90 µm, with repetition rates of 6–8 Hz and an energy density of approximately 6 J/cm2, adjusted based on sample Sr concentration. Sr isotopic data were acquired in static multi-collection mode using Faraday collectors with low resolution. Each analysis involved an initial 30s measurement of the Kr gas blank with the laser off, followed by 60s of data acquisition during ablation. Approximately 4 mL/min of nitrogen gas was added to the carrier gas to improve sensitivity. Data reduction was performed offline. Isobaric interferences were corrected sequentially for Kr, Yb2+, Er2+ and Rb. Finally, the 87Sr/86Sr ratios were calculated and normalized to an interference-corrected 86Sr/88Sr ratio using an exponential law. Data reduction was executed using a custom Excel VBA (Visual Basic for Applications) macro program.

[Apatite] Two apatite reference materials, AP1 and Slyudyanka, were analysed for data quality assessment. The measured 87Sr/86Sr values of ___ ± __ (2s) and ___ ± __ (2s) were consistent with the recommended values of 0.711370 ± 0.000031 (2s) and 0.707683 ± 0.000025 (2s), respectively, within analytical uncertainty (Yang et al., 2014).

[Bastnasite] Two bastnasite reference materials, K-9 and MAD809, were analysed for data quality assessment. The measured 87Sr/86Sr values of ___ ± __ (2s) and ___ ± __ (2s) were consistent with the recommended values of 0.70558 ± 0.00002 (2s) and 0.70845 ± 0.00003 (2s), respectively, within analytical uncertainty (Yang et al., 2019).

[Clinopyroxene] Two clinopyroxene reference materials, YY09-47 and YY09-04, were analysed for data quality assessment. The measured 87Sr/86Sr values of ___ ± __ (2s) and ___ ± __ (2s) were consistent with the recommended values of 0.704052 ± 0.000010 (2s) and 0.703405 ± 0.000010 (2s), respectively, within analytical uncertainty (Zhao et al., 2020).

[Plagioclase] Two plagioclase reference materials, BDL and JH-56, were analysed for data quality assessment. The measured 87Sr/86Sr values of ___ ± __ (2s) were consistent with the recommended values of 0.70533 ± 0.00004 (2s) and 0.71264 ± 0.00015 within analytical uncertainty (Yang et al., 2009b; Xu et al., 2022).

[Calcite] Two calcite reference materials, TLM and LSJ07, were analysed for data quality assessment. The measured 87Sr/86Sr values of ___ ± __ (2s) and ___ ± __ (2s) were consistent with the recommended values of 0.70970 ± 0.00003 (2s) and 0.71004 ± 0.00003 (2s), respectively, within analytical uncertainty (Wu et al., 2023).

[Vesuvianite] Two vesuvianite reference materials, M6635 and M1377, were analysed for data quality assessment. The measured 87Sr/86Sr values of ___ ± __ (2s) and ___ ± __ (2s) were consistent with the recommended values of 0.70826 ± 0.00007 (2s) and 0.70823 ± 0.00006 (2s), respectively, within analytical uncertainty (Wei et al., 2024).

如果数据需要进行外部校正，则上述矿物的两个参考物质中的一个用作校正标样（primary reference material），另一个用作监控标样（monitoring reference material），并添加以下陈述“The uncertainty of the primary reference material was propagated to the final uncertainty of the sample measurements”。

参考文献

Yang, Y., Wu, F., Wilde, S.A., Liu, X., Zhang, Y., Xie, L., Yang, J., 2009a. In situ perovskite Sr-Nd isotopic constraints on the petrogenesis of the Ordovician Mengyin kimberlites in the North China Craton. Chemical Geology 264, 24-42.

Yang, Y.H., Wu, F.Y., Xie, L.W., Yang, J.H., Zhang, Y.B., 2009b. In-situ Sr isotopic measurement of natural geological samples by LA-MC-ICP-MS. Acta Petrologica Sinica 25, 3431-3441.

Yang, Y., Wu, F., Yang, J., Chew, D.M., Xie, L., Chu, Z., Zhang, Y., Huang, C., 2014. Sr and Nd isotopic compositions of apatite reference materials used in U-Th-Pb geochronology. Chemical Geology 385, 35-55.

Yang, Y., Wu, F., Li, Q., Rojas-Agramonte, Y., Yang, J., Li, Y., Ma, Q., Xie, L., Huang, C., Fan, H., Zhao, Z., Xu, C., 2019. In Situ U-Th-Pb Dating and Sr-Nd Isotope Analysis of Bastnasite by LA-(MC)-ICP-MS. Geostandards and Geoanalytical Research 43, 543-565.

Xu, L., Yang, J. H., Wang, H., Li, Y., Zhou, B. Q., Yang, Y. H., Wu, S. T., 2022. A natural plagioclase reference material for microbeam Sr isotopic analysis. Journal of Analytical Atomic Spectrometry, 37(8), 1706-1714.

Wu S, Yang Y, Li T, Huang C, Bao Z, Li Y, Li C, Xu L, Wang H, Xie L, Yang J, Wu F. 2023. Calcite TLM and LSJ07– two natural low-Sr reference materials for microbeam Sr isotope analysis. Journal of Analytical Atomic Spectrometry. 38, 2528-2537.

Wei Q.D., Yang Y.H., Wang H., Wu S.T., Yang M., Huang C., Xu L., Xie L.W., Yang J.H., Wu F.Y., 2024. In situ Sr–Nd isotope analysis of vesuvianite by LA–MC–ICP–MS: methodology and application. Journal of Analytical Atomic Spectrometry, 39,

Zhao H., Zhao X.M., Le Roux P.J., Zhang W., Wang H., Xie L.W., Huang C., Wu S.T., Yang J.H., Wu F.Y., Yang Y.H., 2020. Natural Clinopyroxene reference materials for in situ Sr isotopic measurement by LA-MC-ICP-MS. Frontiers in Chemistry, doi: 10.3389/fchem.2020.594316.

3.3 In situ Nd isotope analysis by LA-MC-ICP-MS

In situ Nd isotopic measurements were performed using a Neptune (Plus) multi-collector inductively coupled plasma mass spectrometer (MC-ICP-MS) (Thermo Fisher Scientific, USA). The analyses were conducted at the Institute of Geology and Geophysics, Chinese Academy of Sciences, Beijing, China (IGGCAS), using a GeoLasPro 193 nm ArF excimer laser (Coherent, Germany) or an Analyte G2 193 nm ArF excimer laser (PhotonMachines, USA) equipped with an ANU HelEx 2-volume cell. Detailed analytical procedures are provided in Yang et al. (2008).如果在分析过程中还用到了X截取锥和氮气增敏技术，还需要增加一下描述“High performance X skimmer cone with the addition of nitrogen was used to increase the sensitivity, enlarge the mass bias stability zone and suppress the nonlinear mass-dependent fractionation (Xu et al., 2015).”

Laser ablation was performed using a spot size ranging from 16 to 60 µm, a repetition rate of 6-8 Hz, and an energy density of ~6 J/cm2, depending on the sample Nd concentration. Each spot analysis involved a 60s data acquisition with the laser fire on. Accurate measurement of 147Sm/144Nd and 143Nd/144Nd ratios by LA-MC-ICP-MS requires careful correction for isobaric interference from 144Sm on the 144Nd signal. This correction is complicated by the fact that the 146Nd/144Nd ratio, typically used for normalizing other Nd isotope ratios, is also affected by Sm interference. Therefore, the mass bias correction for 144Sm interference on 144Nd cannot be directly derived from the measured 146Nd/144Nd ratio. To address this, we employed the revised Sm isotopic abundances (147Sm/149Sm = 1.08680 and 144Sm/149Sm = 0.22332) (Dubois et al., 1992; Isnard et al., 2005). The measured 147Sm/149Sm ratio was first used to calculate the Sm fractionation factor and the measured 147Sm intensity, with the natural 147Sm/144Sm ratio of 4.866559 (calculated from the abovementioned Sm isotopic abundances) applied to estimate the Sm interference on mass 144. The interference-corrected 146Nd/144Nd ratio was then used to determine the Nd fractionation factor. Finally, 143Nd/144Nd and 145Nd/144Nd ratios were normalized using the exponential law. The 147Sm/144Nd ratio of unknown samples was similarly corrected for the 144Sm interference. Raw data were exported offline and processed using an in-house Excel VBA (Visual Basic for Applications) macro program.

[Apatite] Two apatite reference materials, AP1 and AP2, were analysed for data quality assessment. The measured 143Nd/144Nd ratios were ___ ± __ (2s) and ___ ± __ (2s), respectively, consistent with the recommended values of 0.511352 ± 0.000024 (2s) and 0.511007 ± 0.000030 (2s), within analytical uncertainty (Yang et al., 2014). The 147Sm/144Nd ratio was externally calibrated against that of the AP1 reference material (0.0825 ± 0.0012, 2s) during the analytical sessions (Yang et al., 2014), and the uncertainty for the 147Sm/144Nd ratio of AP1 apatite was propagated to the final uncertainty of the sample measurements.

[Monazite] Two monazite reference materials, Manangoutry and M4, were analysed for data quality assessment. The measured 143Nd/144Nd ratios were ___ ± __ (2s) and ___ ± __ (2s), respectively, consistent with the recommended values of 0.511044 ± 0.000022 (2s) and 0.511761 ± 0.000020 (2s), within analytical uncertainty (Liu et al., 2012). The 147Sm/144Nd ratio was externally calibrated against that of the Manangoutry reference material (0.0711 ± 0.0005, 2s) during the analytical sessions (Liu et al., 2012), and the uncertainty for the 147Sm/144Nd ratio of Manangoutry monazite was propagated to the final uncertainty of the sample measurements.

.

[Titanite] Two titanite reference materials, Ontario and BLR-1, were analysed for data quality assessment. The measured 143Nd/144Nd ratios were ___ ± __ (2s) and ___ ± __ (2s), respectively, consistent with the recommended values of 0.512833 ± 0.000008 (2s) and 0.512815 ± 0.000003 (2s), within analytical uncertainty (Xu et al., 2018; Ma et al., 2019). The 147Sm/144Nd ratio was externally calibrated against that of the Ontario reference material (0.1928 ± 0.0011, 2s) during the analytical sessions (Ma et al., 2019), and the uncertainty for the 147Sm/144Nd ratio of Ontario titanite was propagated to the final uncertainty of the sample measurements.

[Bastnasite] Two bastnasite reference materials, K-9 and LZ1384, were analysed for data quality assessment. The measured 143Nd/144Nd ratios were ___ ± __ (2s) and ___ ± __ (2s), respectively, consistent with the recommended values of 0.512767 ± 0.000004 (2s) and 0.512508 ± 0.000010 (2s), within analytical uncertainty (Yang et al., 2019). The 147Sm/144Nd ratio was externally calibrated against that of the K-9 reference material during the analytical sessions (Yang et al., 2019), and the uncertainty for the 147Sm/144Nd ratio of K-9 bastnasite was propagated to the final uncertainty of the sample measurements.

[Allanite] Two allanite reference materials, LE40010 and LE2808, were analysed for data quality assessment. The measured 143Nd/144Nd ratios were ___ ± __ (2s) and ___ ± __ (2s), respectively, consistent with the recommended values of 0.510107 ± 0.000016 (2s) and 0.512230 ± 0.000012 (2s), within analytical uncertainty (Yang et al., 2022). The 147Sm/144Nd ratio was externally calibrated against that of the LE40010 reference material (0.0569 ± 0.0002, 2s) during the analytical sessions (Yang et al., 2022), and the uncertainty for the 147Sm/144Nd ratio of LE40010 allanite was propagated to the final uncertainty of the sample measurements.

[Vesuvianite] Two vesuvianite reference materials, M659 and M6635, were analysed for data quality assessment. The measured 143Nd/144Nd ratios were ___ ± __ (2s), and ___ ± __ (2s), respectively, consistent with the recommended values of 0.512118 ± 0.000012 (2s) and 0.511798 ± 0.000010 (2s), within analytical uncertainty (Wei et al., 2024). The 147Sm/144Nd ratio was externally calibrated against that of the M659 reference material (0.0792 ± 0.0007, 2s) during the analytical sessions (Wei et al., 2024), and the uncertainty for the 147Sm/144Nd ratio of M659 vesuvianite was propagated to the final uncertainty of the sample measurements.

如果Nd同位素比值数据需要进行外部校正，则上述矿物的两个参考物质中的一个用作校正标样（primary reference material），另一个用作监控标样（monitoring reference material），并添加以下陈述“The uncertainty for the 143Nd/144Nd ratios of the primary reference material was propagated to the final uncertainty of the sample measurements”。

参考文献

Dubois, J.C., Retali, G., Cesario, J., 1992. Isotopic analysis of rare earth elements by total vaporization of samples in thermal ionization mass spectrometry. International Journal of Mass Spectrometry and Ion Processes 120, 163-177.

Isnard, H., Brennetot, R., Caussignac, C., Caussignac, N., Chartier, F., 2005. Investigations for determination of Gd and Sm isotopic compositions in spent nuclear fuels samples by MC ICPMS. International Journal of Mass Spectrometry 246, 66-73.

Yang, Y., Wu, F., Wilde, S.A., Liu, X., Zhang, Y., Xie, L., Yang, J., 2009. In situ perovskite Sr-Nd isotopic constraints on the petrogenesis of the Ordovician Mengyin kimberlites in the North China Craton. Chemical Geology 264, 24-42.

Yang, Y., Wu, F., Yang, J., Chew, D.M., Xie, L., Chu, Z., Zhang, Y., Huang, C., 2014. Sr and Nd isotopic compositions of apatite reference materials used in U-Th-Pb geochronology. Chemical Geology 385, 35-55.

Yang, Y., Wu, F., Li, Q., Rojas-Agramonte, Y., Yang, J., Li, Y., Ma, Q., Xie, L., Huang, C., Fan, H., Zhao, Z., Xu, C., 2019. In Situ U-Th-Pb Dating and Sr-Nd Isotope Analysis of Bastnasite by LA-(MC)-ICP-MS. Geostandards and Geoanalytical Research 43, 543-565.

Wei Q.D., Yang Y.H.,* Wang H., Wu S.T., Yang M., Huang C., Xu L., Xie L.W., Yang J.H., Wu F.Y. 2024. In situ Sr–Nd isotope analysis of vesuvianite by LA–MC–ICP–MS: methodology and application. Journal of Analytical Atomic Spectrometry, 39, 1284-1301.

Yang, M., Yang, Y.H.,* Kamo S.L. Romer R. L. Roberts N.M.W. Wang H., Xie L.W., Huang, C., Yang, J.H. Wu, F.Y. 2022. Natural allanite reference materials for in situ U–Th-Pb and Sm–Nd isotopic analysis by LA-(MC)-ICP-MS. Geostandards and Geoanalytical Research,46(2), 169–203.

Ma Q., Evans N.J., Ling X.X., Yang J.H., Wu F.Y., Zhao Z.D., Yang Y.H., 2019. Natural titanite reference materials for In situ U-Pb and Sm-Nd isotopic measurements by LA-(MC)-ICP-MS. Geostandards and Geoanalytical Research. 43, 355-384.

Liu Z. C., Wu, F. Y., Yang, Y. H., Yang, J. H., Wilde, S.A., 2012. Neodymium isotopic compositions of the standard monazites used in U-Th-Pb geochronology. Chemical Geology, 334, 221-239.

Xu, L., Hu, Z.C., Zhang, W., Yang, L., Liu Y., Gao, S., Luo, T., Hu, S.H., 2015. In situ Nd isotope analyses in geological materials with signal enhancement and non-linear mass dependent fractionation reduction using laser ablation MC-ICP-MS, Journal of Analytical Atomic Spectrometry, 30, 232-244.

Xu, L., Yang. J.H., Ni, Q., Yang, Y.H., Hu, Z.C., Liu, Y.S., Wu, Y.B., Luo, T., Hu, S.H., 2018. Determination of Sm-Nd isotopic compositions in fifteen geological materials using laser ablation MC-ICP-MS and application to monazite Geochronology of metasedimentary rock in the North China Craton, Geostandards and Geoanalytical Research, 42, 379-394.

3.4 Sr-Nd-Hf isotope analysis by solution MC-ICP-MS

方法描述

Analytical protocol of Sr-Nd-Hf isotopic compositions by Neptune Plus MC-ICP-MS after DGA and Sr resin chemical purification for one sample digestion. Chemical purifications were undertaken by ion-exchange resin. The analytical protocol is described in detail and can be found elsewhere (Yang et al., 2010, 2011, 2012), and only a brief description is given here.

About 100-150 mg fine sample powder were dissolved in 2.5 ml concentrated HF, 0.2 ml HNO3 and HClO4 on a hot plate (or using steel-jacketed Teflon bombs placed in an oven at 190 °C) for one week. Following complete dissolution, each sample was dried down at high temperature (fuming HClO4) on a hot plate; then the mixture was treated with 14 M HNO3, evaporated to dryness overnight, and taken up in 3 M HNO3 + 3% m/v H3BO3. The capsule was resealed and placed on a hot plate at 100 °C overnight in preparation for chemical purification.

The first stage was the separation of Sr, Nd and Hf from the matrix using Eichrom DGA resin (50–100 μm, 2 ml). The major element fraction were eluted and collected using 3 M HNO3 + 3% m/v H3BO3, meanwhile the Sr fraction was collected for further purification. Then the column was rinsed with 12 M HNO3 to remove effectively any remaining Ca before collection of the Hf fraction, followed by separation of Hf using 3.5 M HNO3+0.2 M HF mixture. Finally, the Nd fraction was eluted with 2 M HCl. The second stage involved the Sr fraction was further purified by Sr-specific resin (100–150 μm, 0.2 ml) prior to mass spectrometric measurement (Yang et al., 2012).

Sr, Nd, and Hf isotopic ratios were measured using a Thermo Scientific Neptune Plus MC-ICP-MS. Whole procedural blanks were less than 100 pg Sr, 50 pg for Nd, and 50 pg for Hf. The 87Sr/86Sr, 143Nd/144Nd and 176Hf/177Hf ratios were normalized to 86Sr/88Sr=0.1194, 146Nd/144Nd=0.7219 and 179Hf/177Hf=0.7325, respectively, using the exponential law. During the period of data acquisition, standard analyses yielded results of 87Sr/86Sr=0.710249±12 (2SD, n=15) for NBS987, 143Nd/144Nd=0.512115±12 (2SD, n=15) for JNdi-1, and 176Hf/177Hf=0.282185±10 (2SD, n=15) for Alfa Hf 14374.

In addition, USGS reference materials BCR-1 and BHVO-2 were also processed for Sr–Nd–Hf isotopes, and gave ratios of 0.704977±12 and 0.703474±11 for 87Sr/86Sr, 0.512655±05 and 0.513000±15 for 143Nd/144Nd and 0.282866±06 and 0.283109±05 for 176Hf/177Hf, respectively, which is identical, within error, to the recommended values (Weis et al., 2005, 2006, 2007).

参考文献

Weis, D., Kieffer, B., Maerschalk, C., Pretorius, W., Barling, J., 2005. High-precision Pb-Sr-Nd-Hf isotopic characterization of USGS BHVO-1 and BHVO-2 reference materials. Geochemistry, Geophysics, Geosystems 6, Q02002, doi:10.1029/2004GC000852.

Weis, D., Kieffer, B., Maerschalk, C., Barling, J., de Jong, J., Williams, G.A., Hanano, D., Pretorius, W., Mattielli, N., Scoates, J.S., Goolaerts, A., Friedman, R.M., Mahoney, J.B., 2006. High-precision isotopic characterization of USGS reference materials by TIMS and MC-ICP-MS. Geochemistry, Geophysics, Geosystems 7, Q08006, doi:10.1029/2006GC001283.

Weis, D., Kieffer, B., Hanano, D., Nobre Silva, I., Barling, J., Pretorius, W., Maerschalk, C., Mattielli, N., 2007. Hf isotope compositions of U.S. Geological Survey reference materials. Geochemistry Geophysics Geosystems 8, Q06006.

Yang, Y.H., Zhang, H.F., Chu, Z.Y., Xie, L.W., Wu, F.Y., 2010. Combined chemical separation of Lu, Hf, Rb, Sr, Sm and Nd from a single rock digest and precise and accurate isotope determinations of Lu-Hf, Rb-Sr and Sm-Nd isotope systems using Multi-Collector ICP-MS and TIMS. International Journal of Mass Spectrometry 290, 120-126.

Yang, Y., Chu, Z., Wu, F., Xie, L., Yang, J., 2011. Precise and accurate determination of Sm, Nd concentrations and Nd isotopic compositions in geological samples by MC-ICP-MS. Journal of Analytical Atomic Spectrometry 26, 1237-1244.

Yang, Y., Wu, F., Liu, Z., Chu, Z., Xie, L., Yang, J., 2012. Evaluation of Sr chemical purification technique for natural geological samples using common cation-exchange and Sr-specific extraction chromatographic resin prior to MC-ICP-MS or TIMS measurement. Journal of Analytical Atomic Spectrometry 27, 516-522.